| Cornell High Energy Synchrotron Source |  |  |  |  |
| --- | --- | --- | --- | --- |
| Doc#: SOP-DET- NNN | Procedure: Use of “multi_pilatus_v4.mac” to run multiple Pilatus detectors from SPEC in various trigger modes, with support for fly-scanning using the USBCTR. |  |  | Prepared by: ARW |
| Rev.: 3 | Date Revised: 8/17/2020 | Date Effective: | Date Expires: | Approved by: |

Purpose

To run one or more Pilatus detectors within SPEC scans, using one of a variety of different trigger modes. For instance:

- Perform “tseries”, “dscan” or “mesh” scans where a snapshot is obtained for each point in the scan, for each configured detector, and saved to a directory on CHESS_DAQ.

- Perform “pil_movie <Nimages> <dwell>” to initiate an image series from one more pilatus detectors.

- Perform “flyscan” or “flymesh” to obtain high frame-rate images from one or more Pilatus detectors. (scans provided by flyscan_v4.mac)

Equipment and Setup Needed

- PILATUS Pixel Array Detector(s) and associated power supplies, dry N2, chiller (for Pilatus 300K and above). See PILATUS hardware SOPs for details, but note that the macros described here supersede the SPEC macros described in those documents (as of August 2020).

- A working EPICS IOC for each PILATUS.

- For saving images to CHESS_DAQ: Station-specific CHESS_DAQ location mounted on station computer. For example, “/mnt/currentdaq” on the PILATUS control computer should point to  “/nfs/chess/daq/20NN-N/id3b”. These macros also permit saving images to the Pilatus control computer.

- “STATION_ID” global variable defined within SPEC to match the string appearing in CHESS_DAQ, e.g. “id3b”. (Best to assign in userlist.mac or station.mac)

- Trigger-mode-specific wiring for trigger modes other than “Internal”.

- Macro set loaded into SPEC. For example if the macro file “multi_pilatus_v4.mac”. is located in the “Macros” directory of the station  home directory, then issue “udo multi_pilatus_v4.mac”

Safety – see hardware specific SOPs.

Overview

Sections A-C below provide several example command sequences for particular setups, each sharing the same minimal command sequence: pil_setup and pil_on. The additional commands pil_setdir and pil_settrig are required to change the target directory and/or trigger mode, and pil_von/voff is provided to lend command line access to video modes. The complete function list is provided in Section D, and Section E gives an example SPEC session, showing macro output and the resulting image files resulting from 1D and 2D scans.

Procedure

NOTE: Ensure detector window is covered with the protective plate when not in use to prevent damage from exposure.

A. Example  1:  one detector with internal triggering (the 300K as “PIL5”)

| Command | Action |
| --- | --- |
| pil_setup PIL5 | Sets up the detector with EPICS prefix “PIL5” as Pilatus number “0”, to save images to the current directory on CHESS_DAQ. |
| pil_settrig “Internal” | Selects Internal trigger mode use when taking snapshots within scans. NOTE: this has no effect on the TriggerMode EPICS PV until a scan is performed. |
| pil_on | Turns all detectors (in this case, only PIL5) “on”, such that a spec scan will trigger one snapshot per scan point. |
| pil_movie 10 .1 | Immediately take 10 images, 0.1-second each. (no metadata). |
| tseries 5 .5 | Perform a five-point time series, taking a half-second snapshot at each point. Images will have the format: <specfile_PIL5_SCAN#_POINT#> |
| dscan samx 0 5 1 | Perform a single motor scan, taking a 1-second snapshot at each point. Images have the same format as above. |
| mesh samx 0 1 samz 0 1 1 | Perform a 2D, 2-motor scan, taking 1 second snapshots at each. point. Images will have the format: <specfile_PIL5_SCAN#_ROW#_POINT#> |
| pil_off | Turn off snapshot acquisition. |
| pil_unsetup all | Erase SPEC’s knowledge of PILATUS detectors and associated macros. (This amounts to erasing the associative array “PIL”) |

B. Example 2: two detectors and external triggering ( “PIL5” and “PIL9”)

| Command | Action |
| --- | --- |
| pil_setup PIL5 0 PIL5 or pil_setup PIL5 PIL5 | Sets up the detector with EPICS prefix “PIL5” as Pilatus number “0”, to save images to the “PIL5” subdirectory of the current directory on CHESS_DAQ. |
| pil_setup PIL9 1 PIL9 | As above, but sets up the detector “PIL9” as Pilatus number “1”, to save images to the “PIL9” subdirectory |
| pil_settrig “Mult. Trigger” | Selects Multi trigger mode on both detectors when taking snapshots within scans. This implies that there is digital output / hardware trigger hooked into both detectors, that will go “high” for each point in any scan. |
| pil_on 0 1 or pil_on | Enable in-scan snapshots for both detectors. |
| “tseries”, “dscan”, and “mesh” | Perform scans as in Example 1. Images will have the format <specfile_PIL#_SCAN#_POINT#> or <specfile_PIL#_SCAN#_ROW#_POINT#> for 1D or 2D scans, respectively. |
| pil_off 0 | Turn off detector 0, such that only PIL9 will take images. |

C. Example 2: three detectors in fly mode ( “PIL5”, “PIL9” and “PIL6”)

In addition to “multi_pilatus_v4.mac”, this functionality requires “flyscan_v4.mac”

| Command | Action |
| --- | --- |
| pil_setup PIL5 | Sets up the detector with EPICS prefix “PIL5” (the 300K) as Pilatus number “0”, to save images to the current CHESS_DAQ directory. |
| pil_setup PIL9 1 | As above, but sets up the detector “PIL9” (a 200K) as Pilatus number “1”. |
| pil_setup PIL6 2 | As above, but sets up the detector “PIL6” (a 100K) as Pilatus number “2”. |
| pil_settrig “Mult. Trigger” | Selects Multi trigger mode on both detectors when taking snapshots within scans. For fly scanning (below), this requires that the USBCTR device is setup as described in flyscan_v4.mac, with step signal from the MDCP going into it, and the “CLK1” output going to the detectors. Also: a digital I/O output (see FLY_TRIGCHAN in id3b_NI_DIO.mac ~/Macros) goes to the USBCTR. |
| pil_on 0 1 2 or pil_on | Enable in-scan snapshots for all detectors. |
| flysetup | Set up the fly mode fast and slow motors, and verify USBCTR setup. |
| flyscan 10 .5 100 | Scan the fast flymode motor 10 mm, taking 0.1 second. images every 0.5 mm (20 images). Images have the format <specfile_PIL#_SCAN#_POINT#>. |
| flymesh 10 .1 2 .1 50 | Perform a 10 mm x 2 mm 2D flymode map, with 0.1 mm pitch and 50 miliseconds per point (100 x 20 point map). Images have the format <specfile_PIL#_SCAN#_ROW#_POINT#>. |

D. Complete Function List and Usages (taken from mult_pilatus_v4.mac).

# pil_setup PREFIX [N] [path_arg] / pil_unsetup n1 [n2 ..] | all

#     pil_setup: if N is omitted, use N=0

#       e.g. pil_setup PIL5 (set up PIL5 as the 0th pilatus, and set directory to CWD on DAQ)

# pil_unsetup: Un-setup one or more detectors, explicitly designated by number of “all” for all.

#       e.g  pil_unsetup 0 1 (unsetup detectors 0 and 1)

#	     pil_unsetup all (unsetup all detectors and delete the associative array “PIL”)

# pil_setdir [n1 [n2 ..]]  [newpath]

# pil_settrig [n1 [n2 ...]] trigmode

#	Set trigger mode (within spec scans) for 1 or more detectors for 1 or more detectors:

#	"Internal", "Ext. Enable","Mult. Trigger", or "Ext. Trigger"

# pil_show:

#     Lists Pilati that are set up and whether they are active, etc. Values from PIL

# pil_prout:

#     Lists detail (from EPICS PVs) for active detectors, including E-thresh.

# pil_on/off

#	set target path for one or more detectors that are already set up

#	if newpath is -1 or absent, use current directory, which MUST be on the DAQ

# pil_von/voff

#	sets all active detectors  to "Alignment" mode and begins taking snapshots.

# pil_movie  Nimages  dwell

# 	Sets up and initiates a series of pilatus images by all detectors that are on. If

#	the trigger mode is "Internal", then the movie begins immediately. Otherwise the detector

#		will await one or more triggers (depending on mode)

# pil_im_get:

#      Usage: pil_im_get(n, pv_suffix [, "string"]) (a standin for e.g.

#	epics_get("PIL5:cam1:<pv_suffix>" [, "string"])

# pil_im_set:

#	  Usage: pil_im_set(n, pv_suffix, val)

#			 (a standin for e.g. epics_put("PIL5:cam1:<pv_suffix>", val)

# pil_get/set: Like pil_im_get/set but where  "cam1" is absent from the PV string:

#			 (a standin for e.g. epics_get("PIL5:<pv_suffix>" [, "string"])

# _pil_scan_head: A pilatus-specific commands to be hooked to “scan_head” in the standard scan

#	sequence. Determines whether Trigger Mode, scan dimensionality, and if flymode is active.

#	Note that the file path is set here, but filename string is set in _pil_precount, since

#	for 2D scans it must be reset for each row.

# _pil_precount: A hook for user_precount macro that determines when to set the filename and

#   	initiate acquisition.

# _pil_file_setpath, _pil_file_setname:

#		sets filename based on scan dimensionality (and possibly trigger mode)

E. Example SPEC session

2439.SPEC> pwd

/mnt/home/chess_id3b

2440.SPEC> pil_setup  PIL5

Error: pil_get_path failed to locate current directory or find STATION_ID id3b on DAQ

STATION_ID should be set in userlist.mac or station.mac, e.g.

Error: In pil_setdir: cannot determine path.

If valid, current path will be retained

pil_setup: Error determining correct path -- aborting

0 detectors defined

2441.SPEC> cd  /mnt/currentdaq/id3b-misc-1/

Now in "/nfs/chess/daq/2020-2/id3b/id3b-misc-1".

2442.SPEC> pil_setup  PIL5

1 detectors defined

Pilatus 0/PIL5 is OFF; Mode:      Internal; path:/mnt/currentdaq/id3b-misc-1

2443.SPEC> pil_on

1 detectors defined

Pilatus 0/PIL5 is  ON; Mode:      Internal; path:/mnt/currentdaq/id3b-misc-1

2444.SPEC> newfile

Data file (testpil4)? testpil5

Last scan number (0)?

Using "testpil5".  Next scan is number 1.

2445.SPEC> ls

2446.SPEC> tseries 5 1

Total 5 points, 6 seconds

Scan   1   Mon Aug 17 18:33:11 2020   file = testpil5  spec  user = chess_id3b

tseries 5 1 0

#      Time     Seconds         ic1    Seconds         ic2        azer        bedh       diode    VBPM VER    VBPM HOR

1 detectors defined

Pilatus 0/PIL5 is  ON; Mode:      Internal; path:/mnt/currentdaq/id3b-misc-1

0    0.0000           1         110          1         317        1382           0           0           0           0

1    1.3421           1         110          1         318        1385           0           0           0           0

2    2.6180           1         110          1         317        1381           0           0           0           0

3    3.8892           1         109          1         317        1384           0           0           0           0

4    5.1671           1         109          1         316        1380           0           0           0           0

2447.SPEC> pil_setup  PIL5 1

2 detectors defined

Pilatus 0/PIL5 is  ON; Mode:      Internal; path:/mnt/currentdaq/id3b-misc-1

Pilatus 1/PIL5 is OFF; Mode:      Internal; path:/mnt/currentdaq/id3b-misc-1

2448.SPEC> ls

2449.SPEC> pil_off

2 detectors defined

Pilatus 0/PIL5 is OFF; Mode:      Internal; path:/mnt/currentdaq/id3b-misc-1

Pilatus 1/PIL5 is OFF; Mode:      Internal; path:/mnt/currentdaq/id3b-misc-1

2450.SPEC> pil_on  1

2 detectors defined

Pilatus 0/PIL5 is OFF; Mode:      Internal; path:/mnt/currentdaq/id3b-misc-1

Pilatus 1/PIL5 is  ON; Mode:      Internal; path:/mnt/currentdaq/id3b-misc-1

2451.SPEC> wm xrfx  xrfz

xrfx       xrfz

xrfx       xrfz

User

High      139.92500  150.00000

Current     1.00000   22.00000

Low       -60.07500 -130.00000

Dial

High       57.03550   67.03550

Current   -81.88950  -60.96450

Low      -142.96450 -212.96450

2452.SPEC> mesh  xrfx 0 1 2 xrfz 21  22 2 .5

Total 9 points, 6.3 seconds

Scan   2   Mon Aug 17 18:34:45 2020   file = testpil5  spec  user = chess_id3b

mesh xrfx 0 1 2 xrfz 21 22 2  0.5

#      xrfx      xrfz     Seconds         ic1    Seconds         ic2        azer        bedh       diode    VBPM VER    VBPM HOR

2 detectors defined

Pilatus 0/PIL5 is OFF; Mode:      Internal; path:/mnt/currentdaq/id3b-misc-1

Pilatus 1/PIL5 is  ON; Mode:      Internal; path:/mnt/currentdaq/id3b-misc-1

Inside _pil_file_setname: n_slow = 0

0    0.0000   21.0000         0.5          58        0.5         160         691           0           0           0           0

1    0.5000   21.0000         0.5          58        0.5         160         691           0           0           0           0

2    1.0000   21.0000         0.5          58        0.5         160         691           0           0           0           0

Inside _pil_file_setname: n_slow = 1

3    0.0000   21.5000         0.5          58        0.5         160         690           0           0           0           0

4    0.5000   21.5000         0.5          58        0.5         161         692           0           0           0           0

5    1.0000   21.5000         0.5          59        0.5         160         692           0           0           0           0

Inside _pil_file_setname: n_slow = 2

6    0.0000   22.0000         0.5          58        0.5         160         691           0           0           0           0

7    0.5000   22.0000         0.5          58        0.5         160         692           0           0           0           0

8    1.0000   22.0000         0.5          57        0.5         161         691           0           0           0           0

2453.SPEC> ls

2454.SPEC> pil_off

2 detectors defined

Pilatus 0/PIL5 is OFF; Mode:      Internal; path:/mnt/currentdaq/id3b-misc-1

Pilatus 1/PIL5 is OFF; Mode:      Internal; path:/mnt/currentdaq/id3b-misc-1

2455.SPEC> pil_settrig "Mult. Trigger"

2 detectors defined

Pilatus 0/PIL5 is OFF; Mode: Mult. Trigger; path:/mnt/currentdaq/id3b-misc-1

Pilatus 1/PIL5 is OFF; Mode: Mult. Trigger; path:/mnt/currentdaq/id3b-misc-1

2456.SPEC> pil_unsetup  1

1 detectors defined

Pilatus 0/PIL5 is OFF; Mode: Mult. Trigger; path:/mnt/currentdaq/id3b-misc-1

2457.SPEC> pil_on

1 detectors defined

Pilatus 0/PIL5 is  ON; Mode: Mult. Trigger; path:/mnt/currentdaq/id3b-misc-1

2458.SPEC> flymesh 5 1 2 1 500

Estimated length of scan in seconds / minutes / hours : 11 / 0.2 / 0.00

Scan Origin is x=1.00, y=22.00

x pitch: 1

y pitch: 1

x scan size: 5

y scan size: 2

dwell: 500

Fast axis speed is 8000 Hz, 2.00 mm/sec

Normal speed of xrfx is 50000 Hz = 12.50 mm / sec

Desired scan speed is 8000 Hz = 2.00 mm sec

Total 3 points, 2.1 seconds

Scan   3   Mon Aug 17 18:36:23 2020   file = testpil5  spec  user = chess_id3b

flymesh  xrfx 1.5 5.5 4 xrfz 22 24  2 0.5

#      xrfz  Seconds      ic1    Seconds      ic2     azer     bedh    diode VBPM VER VBPM HOR

Using 3 counters from "USBCTR2:MCS".

1 detectors defined

Pilatus 0/PIL5 is  ON; Mode: Mult. Trigger; path:/mnt/currentdaq/id3b-misc-1

DEBUG: Beginning _scan_on deff

Restoring xrfx velocity to 50000 Hz

before move

after move

Normal speed of xrfx is 50000 Hz = 12.50 mm / sec

Desired scan speed is 8000 Hz = 2.00 mm sec

Begin row 1 of 2; raster xrfx to 6.000

In fly_precount -- about to send EraseStart

Inside _pil_file_setname: n_slow = 0

In fly_getcounts -- about to send stopall

0   22.0000      2.5      283        2.5      813     3462        0        0        0        0

Restoring xrfx velocity to 50000 Hz

before move

after move

Normal speed of xrfx is 50000 Hz = 12.50 mm / sec

Desired scan speed is 8000 Hz = 2.00 mm sec

Begin row 2 of 2; raster xrfx to 1.000

In fly_precount -- about to send EraseStart

Inside _pil_file_setname: n_slow = 1

In fly_getcounts -- about to send stopall

1   23.0000      2.5      283        2.5      804     3457        0        0        0        0

Restoring xrfx velocity to 50000 Hz

before move

after move

Normal speed of xrfx is 50000 Hz = 12.50 mm / sec

Desired scan speed is 8000 Hz = 2.00 mm sec

Begin row 3 of 2; raster xrfx to 6.000

In fly_precount -- about to send EraseStart

Inside _pil_file_setname: n_slow = 2

In fly_getcounts -- about to send stopall

2   24.0000      2.5      284        2.5      810     3455        0        0        0        0

Restoring xrfx velocity to 50000 Hz

Returning to xrfx = 1.00, xrfz =22.00

Scan time elapsed in seconds / minutes / hours : 13 / 0.2 / 0.00

2459.SPEC> ls

1				 testpil5_PIL5_002_000_0001.tiff  testpil5_PIL5_003_000_0004.tiff

crl_align			 testpil5_PIL5_002_000_0002.tiff  testpil5_PIL5_003_001_0000.log

n				 testpil5_PIL5_002_001_0000.tiff  testpil5_PIL5_003_001_0000.tiff

test				 testpil5_PIL5_002_001_0001.tiff  testpil5_PIL5_003_001_0001.tiff

testdir				 testpil5_PIL5_002_001_0002.tiff  testpil5_PIL5_003_001_0002.tiff

testpil1			 testpil5_PIL5_002_002_0000.tiff  testpil5_PIL5_003_001_0003.tiff

testpil5			 testpil5_PIL5_002_002_0001.tiff  testpil5_PIL5_003_001_0004.tiff

testpil5_PIL5_001_0000.tiff	 testpil5_PIL5_002_002_0002.tiff  testpil5_PIL5_003_002_0000.log

testpil5_PIL5_001_0001.tiff	 testpil5_PIL5_003_000_0000.log   testpil5_PIL5_003_002_0000.tiff

testpil5_PIL5_001_0002.tiff	 testpil5_PIL5_003_000_0000.tiff  testpil5_PIL5_003_002_0001.tiff

testpil5_PIL5_001_0003.tiff	 testpil5_PIL5_003_000_0001.tiff  testpil5_PIL5_003_002_0002.tiff

testpil5_PIL5_001_0004.tiff	 testpil5_PIL5_003_000_0002.tiff  testpil5_PIL5_003_002_0003.tiff

testpil5_PIL5_002_000_0000.tiff  testpil5_PIL5_003_000_0003.tiff  testpil5_PIL5_003_002_0004.tiff

2460.SPEC> pil_off

1 detectors defined

Pilatus 0/PIL5 is OFF; Mode: Mult. Trigger; path:/mnt/currentdaq/id3b-misc-1

F. Trigger Wiring

The following wiring diagrams correspond to  two different   ways to use the “Mult. Trigger” mode with the Pilatus detectors. In the first case, SPEC supplies a single trigger to the USBCTR, just before beginning fly-scan motor motion. (This is currently  done by one of the digital I/O lines on the  National Instruments counter-time board). Subsequently, the output from scalar channel 7  on the USBCTR triggers  frame advances on the Pilatus,  as well as advancing channels for its own multi-channel scalars (used to collect ion chamber signal during each frame).

In the 2nd case, Pilatus triggering is performed with the “output” signal  from the “standard” CHESS counter/timer board. NB: Often this signal is buffered before being sent to the Pilatus.